## Overview
Overview of Inhaler Common Use Errors Matrix
Version 1.0
IPAC-RS Devices Subgroup, 2020
Purpose
This Inhaler Common Use Errors Matrix, a list of use errors that are common to most or many inhalers of each device class, has been complied in order to help companies conduct use-related risk assessments for their inhaled products and, most importantly, to improve consistency between different companies and their products. Similarly, this matrix may assist regulatory authorities in taking a consistent view when evaluating the risk posed by these use errors across different products. Common use errors for the main classes of inhaler are listed along with scores reflecting the likely impact on delivery to the lung of an individual dose. This error matrix is intended to be used as an input to a use related risk assessment for an Inhalation Device or Drug Product. This list has been compiled by the IPAC-RS Common Use Errors Group and cross-checked with relevant literature.
Scope
The scope of this error matrix is to identify use errors which create hazardous situations related to misdosing and is intended to provide a starting point for teams to evaluate the harm resulting from misdosing of their specific drug product. In this way, this use error matrix may provide a starting point for teams seeking to conduct a use related risk assessment specific to their product and ensure some consistency of approach between across the industry and regulators. ISO EN ISO 14971:2019 gives definitions for a hazard, a hazardous situation that may result from a hazard, and harm that may be caused by these, see diagram adapted from C1 right. Here, a use error is a sequence of events that creates a hazardous situation: a change in the dose delivered to the patient lung resulting from the use error.
Harm to the patient as a result of misdosing arising from these use errors is not evaluated here and should be evaluated separately for each individual device or product. Patient outcomes resulting from misdosing are highly dependent on the drug, indication and patient profile. The occurrence rate, or likelihood, of each error is not evaluated here and should be evaluated as part of the use related risk assessment and human factors programme, for example supported by user study data where available. Likewise, the effect of use error repetition at a given frequency of occurrence, and whether that frequency mitigates or poses an increased patient risk, should similarly be accounted for as part the product risk assessment. Where an error does not impact dosing but does present another hazardous situation, e.g. potential for microbial infection if not cleaned regularly, then this hazardous situation is listed but not scored. It is nonetheless important that the project team evaluate these hazardous situations appropriately for their product. References ISO 14971:2019 Medical devices – Application of risk management to medical devices
Device Classes in Scope
Pressurised Metered Dose Inhaler (pMDI)
Breath-Actuated Metered Dose Inhaler (BAMDI)
Soft Mist Inhaler (SMI)
Dry Powder Inhaler (DPI) - Manually-Loaded
Dry Powder Inhaler (DPI) - Pre-Loaded
Probable Impact of Inhaler Use Errors on Dose Delivery
Score for most likely effect on delivery to the lung of individual dose.
| Category | Score | Description
| Maximal Effect | 10 | Patient certain to receive no dose, or certain to receive additional doses.
| High Effect | 7 | Reasonable probability of severely reduced or increased dose.
| Moderate Effect | 4 | Reasonable probability of moderate change in dose.
| Minor Effect | 1 | Most likely to have a low impact on dosing.
Use Error
Change in dose
delivered to lung
Compromised medical care
P1 – Probability of Use Error Occurring
P2 – Probable Effect of Use Error on Lung Dose
P3 – Probability of Lung Dose Change Resulting in Patient Harm
Severity of harm
Scope of Common Use Errors Matrix
Scope of Product Use-Related Risk Assessment
P1 x P2 x P3 x Severity = Risk
Device User Interface
Scope of Common Use Errors Matrix and Relationship Between Hazard, Sequence of Events Hazardous Situation and Harm - Adapted from ISO 14971:2019
Figure C.1. ISO 14971:2019
Definitions for Common Use Errors Matrix

## Common Use Errors Matrix
Inhaler Common Use Errors Matrix
Version 1.0
IPAC-RS Devices Subgroup, 2020
The inclusion of a particular error for a device category does not necessarily imply its presence in any particular product. Many devices have design features that negate certain errors. This list should be modified and adapted accordingly by project teams.
| Task Step / Sub-Step |  | Associated Use Error | pMDI | BAMDI | SMI | DPI (Manual-Load) | DPI (Pre-Loaded) | Lung Delivery Effect | Comment | Score for Effect on Lung Delivery
1. Management and Care of Inhaler
1.1. Maintainance and Storage of Inhaler
1.1.1 Maintainance and Storage of Inhaler (All Inhalers)
| a. | Follow inhaler maintenance and cleaning routines as instructed. | Device not cleaned regularly as instructed | X | X | X | X | X | Potential for moderate dose reduction. | Additional hazards include, inhalation of foreign object and infection. | 4
| b. | Storage of inhaler at appropriate temperature and humidity conditions | Inhaler not stored at correct temperature or humidity | X | X | X | X | X | Potential for increased or decrased dose | Drug degradation and changes in fine particle mass. Degradants also pose a direct safety hazard. | Not Scored, should be scored by project team.
1.2 Correct Identification of Inhaler for Dose Administration
1.2.1 Correct Identification of Inhaler (All Inhalers)
| a. | Identification of inhaler expiry date | Use of inhaler past shelf life of manufacturered product | X | X | X | X | X | Potential for increased or decrased dose | Drug degradation and changes in fine particle mass. Degradants also pose a direct safety hazard. | Not Scored, should be scored by project team.
| b. |  | Use of inhaler past in-use period (after opening secondary protective pack). | X | X | X | X | X | Potential for increased or decrased dose | Drug degradation and changes in fine particle mass. Degradants also pose a direct safety hazard. | Not Scored, should be scored by project team.
| c. | Identification/differentiation of correct inhaler (e.g. drug, dose strength) | Use of one medicine in place of another (e.g. 'reliever' in place of 'preventer'). | X | X | X | X | X | No dose of correct medication. Inceased dose of other APIs. |  | 10
| d. |  | Use of inhaler containing correct medicine at incorrect strength (e.g. from previous prescription or belonging to a family member). | X | X | X | X | X | Potential for severely increased or reduced dose. |  | 7
| e. | Identification of whether inhaler is empty or not. | Use of inhaler with a zero reading on dose counter but nonetheless achieving an apparently-normal spray. | X | X | X | X | X | Potential for severely reduced dose. | Reduced likelihood for devices with lockout | 7
| f. |  | Use of empty/nearly empty inhaler. | X | X | X | X | X | No dose / severely reduced dose | Reduced likelihood for devices with lockout | 10
| g. | Identification of whether inhaler is functioning correctly. | Use of inhaler with clearly apparent technical issue preventing / severaly limiting dosing. | X | X | X | X | X | Potential for no dose / severely reduced dose. |  | 10 / 7
1.3 Preparation of a New Inhaler
1.3.1 Preparation of a New Inhaler (pMDI)
| a. | New inhaler shaken prior to priming | New inhaler not shaken prior to priming | X |  |  |  |  | Potential for increased or reduced dose for first 1-3 actuations | Suspensions only | 4
| b. | Correct priming of new inhaler | New inhaler not correctly primed | X |  |  |  |  | No dose/reduced dose for first 1-3 actuations | Need dependant on valve type | 4
1.3.2 Preparation of a New Inhaler (BAMDI)
| a. | New inhaler shaken prior to priming | New inhaler not shaken prior to priming |  | X |  |  |  | Potential for increased or reduced dose for first 1-3 actuations | Suspensions only | 4
| b. | Correct priming of new inhaler | New inhaler not correctly primed |  | X |  |  |  | No dose/reduced dose for first 1-3 actuations | Need dependant on valve type | 4
1.3.3 Preparation of a New Inhaler (SMI)
| a. | Correct device assembly, e.g. cartridge or mouthpiece. | Not all parts correctly assembled prior to priming. |  |  | X |  |  | No dose / severely reduced dose |  | 10 / 7
| b. | Correct priming of new inhaler | New inhaler not correctly primed |  |  | X |  |  | No dose/reduced dose for first 1-3 actuations |  | 4
2. Preparation of a New Unit Dose
2.1 Preparation for Dose Administration
2.1.1 Preparation for Dose Administration (All Inhalers)
| a. | Remove inhaler cap/cover | Cap / cover not removed | X | X | X | X | X | No dose |  | 10
| b. | Check for dirt/debris in mouthpiece | No check for dirt/debris in mouthpiece. | X | X | X | X | X | No impact | Inhalation of foreign object and infection hazards. | Not Scored, should be scored by project team.
| c. | Check for blockages in mouthpiece | No check for blockages in mouthpiece. | X | X | X | X | X | Potential for severely reduced dose |  | 7
| d. | Retain cap after removal | Cap discarded before inhaler is empty. | X | X | X | X | X | No impact | Inhalation of foreign object and infection hazards. | Not Scored, should be scored by project team.
2.1.2 Preparation for Dose Administration (pMDI)
| a. | Leave canister in acutator | Canister removed from actuator prior to use | X |  |  |  |  | No dose |  | 10
| b. | Prime inhaler after specified re-priming period (e.g. 5 days) | Inhaler not primed after specified re-priming period (e.g. 5 days) | X |  |  |  |  | Potential for no dose/reduced dose for first 1-3 actuations |  | 4
| c. | Shake inhaler correctly prior to dosing (for specified duration) | Inhaler not correctly shaken (for specified duration) prior to dosing | X |  |  |  |  | Potential for increased or reduced dose. | Suspensions only. Magnitude of dose effect and impactful time period both product dependent. | 7
| d. | Hold device in correct orientation during dose preparation | Device held in incorrect orientation during dose preparation (re-priming, powder exposure) | X |  |  |  |  | Potential for severely reduced dose |  | 7
2.1.3 Preparation for Dose Administration (BAMDI)
| a. | Leave canister in acutator | Canister removed from actuator prior to use |  | X |  |  |  | No dose |  | 10
| b. | Prime inhaler after specified re-priming period (e.g. 5 days) | Inhaler not primed after specified re-priming period (e.g. 5 days) |  | X |  |  |  | Potential for no dose/reduced dose for first 1-3 actuations |  | 4
| c. | Actuate lever (if present) | Lever not actuated or only partially actuated. |  | X |  |  |  | No dose |  | 10
| d. | Shake inhaler correctly prior to dosing (for specified duration) | Inhaler not correctly shaken (for specified duration) prior to dosing |  | X |  |  |  | Potential for increased or reduced dose. | Suspensions only. Magnitude of dose effect and impactful time period both product dependent. | 7
| e. | Hold device in correct orientation during dose preparation | Device held in incorrect orientation during dose preparation (re-priming, powder exposure) |  | X |  |  |  | Potential for severely reduced dose |  | 7
2.1.4 Preparation for Dose Administration (SMI)
| a. | Prime inhaler after specified re-priming period (e.g. 5 days) | Inhaler not primed after specified re-priming period (e.g. 5 days) |  |  | X |  |  | Potential for no dose/reduced dose for first 1-3 actuations |  | 4
| b. | Load inhaler (by turning the device) | Does not load inhaler |  |  | X |  |  | No Dose |  | 10
2.1.5 Preparation for Dose Administration (Manually-Loaded DPIs e.g. Capsules, Blister Disks)
| a. | Open inhaler | Inhaler not opened |  |  |  | X |  | No dose |  | 10
| b. | Load a new/unused capsule/blister into inhaler | Capsule/blister not correctly loaded into inhaler |  |  |  | X |  | No dose |  | 10
| c. | Actuate mechanism to fully pierce capsule/blister prior to inhalation | Piercing mechanism not actuated |  |  |  | X |  | No dose |  | 10
| d. |  | Capsule/blister not fully pierced |  |  |  | X |  | Potential for severely reduced dose |  | 7
| e. |  | Actuates piercing mechanism during or at end of inhalation |  |  |  | X |  | No dose or severly reduced dose |  | 10 / 7
| f. | Ensure device not shaken while powder dose is exposed | Device shaken after exposing powder dose |  |  |  | X |  | Potential for reduced dose | Effect is device dependent. Lower impact for pierced capsules/blisters than "peeled" blister strips | 4
| g. | Hold device in correct orientation during dose preparation | Device held in incorrect orientation during dose preparation (re-priming, powder exposure) |  |  |  | X |  | Potential for reduced dose | Effect is device dependent. Lower impact for pierced capsules/blisters than "peeled" blister strips | 4
2.1.6 Preparation for Dose Administration (Pre-Loaded DPIs e.g. Resevoir, Blister Strip)
| a. | Open inhaler fully | Inhaler not opened |  |  |  |  | X | No dose |  | 10
| b. |  | Inhaler partially opened |  |  |  |  | X | Potential for no dose or severely reduced dose |  | 10 / 7
| c. | Actuate dose loading mechanism (lever, twist) | Dose loading mechanism not actuated |  |  |  |  | X | No dose |  | 10
| d. |  | Dose loading mechanism partially actuated |  |  |  |  | X | Potential for no dose or severely reduced dose |  | 10 / 7
| e. | Ensure device not shaken while powder dose is exposed | Device shaken after/during exposure of powder dose |  |  |  |  | X | Potential for severely reduced dose | Effect is device dependent. Lower impact for pierced capsules/blisters than "peeled" blister strips | 7
| f. | Hold device in correct orientation during dose preparation | Device held in incorrect orientation during dose preparation (re-priming, powder exposure) |  |  |  |  | X | Potential for severely reduced dose | Effect is device dependent. Lower impact for pierced capsules/blisters than "peeled" blister strips | 7
3. Administration of a Single Unit Dose
3.1 Prior to Inhaling
3.1.1. Prior to Inhaling (All Inhalers)
| a. | Exhale fully, as far as comfortable, immediately prior to dose inhalation | No exhalation or partial exhalation immediately prior to dose inhalation | X | X | X | X | X | Potential for reduced dose to small / alveolated airways |  | 4
| b. | Tilt head backwards | Head not tilted backwards to dose | X | X | X | X | X | Low likelihood of reduced dose |  | 1
| c. | Correctly position inhaler mouthpiece in mouth | Inhaler mouthpiece not placed in mouth or incorrectly positioned | X | X | X | X | X | No dose / potential for severely reduced dose |  | 10 / 7
| d. | Ensure good lip seal around mouthpiece | Good lip seal with mouthpiece not achieved | X | X | X | X | X | Potential for reduced dose |  | 4
3.1.2 Prior to Inhaling (pMDI)
| a. | Exhale away from device | Exhalation directed fully or partially into the device | X |  |  |  |  | Potential for reduced dose | Humidity may increase drug deposition on valve. | 4
| b. | Hold device in correct orientation | Device held in incorrect orientation | X |  |  |  |  | Low likelihood of reduced dose |  | 1
3.1.3 Prior to Inhaling (BAMDI)
| a. | Exhale away from device | Exhalation directed fully or partially into the device |  | X |  |  |  | Potential for reduced dose | Humidity may increase drug deposition on valve. | 4
| b. | Hold device in correct orientation | Device held in incorrect orientation |  | X |  |  |  | Low likelihood of reduced dose |  | 1
3.1.4 Prior to Inhaling (SMI)
| a. | Exhale away from device | Exhalation directed fully or partially into the device |  |  | X |  |  | Low likelihood of reduced dose |  | 1
| b. | Hold device in correct orientation | Device held in incorrect orientation |  |  | X |  |  | Low likelihood of reduced dose |  | 1
Prior to Inhaling (Manually-Loaded DPIs)
| a. | Exhale away from device | Exhalation directed fully or partially into the device |  |  |  |  | X | Potential for severely reduced dose |  | 7
| b. | Hold device in correct orientation | Device held in incorrect orientation |  |  |  |  | X | Potential for severely reduced dose |  | 7
Prior to Inhaling (Pre-Loaded DPIs)
| a. | Exhale away from device | Exhalation directed fully or partially into the device |  |  |  |  | X | Potential for severely reduced dose |  | 7
| b. | Hold device in correct orientation | Device held in incorrect orientation |  |  |  |  | X | Potential for severely reduced dose |  | 7
3.2 Inhale Loaded Unit Dose
3.2.1 Inhale Loaded Unit Dose (All Inhalers)
| a. | Inhale through mouth (not nose) | Inhalation through nose and not mouth | X | X | X | X | X | No dose |  | 10
| b. |  | Inhalation through mouth and nose simultaneously | X | X | X | X | X | Potential for reduced dose |  | 4
| c. | Inhale deeply to fill lungs | Incomplete or shallow inhalation | X | X | X | X | X | Potential for reduced dose to small / alveolated airways |  | 4
| d. | Hold breath after dose inhalation | No breath hold following inhalation | X | X | X | X | X | Potential for reduced dose, in particular to small / alveolated airways |  | 4
3.2.2 Inhale Loaded Unit Dose (pMDI)
| a. | Take dose immediately following shaking device | Delay between shaking device and taking dose | X |  |  |  |  | Potential for reduced or increased dose | Effect is formulation dependent | 7
| b. | Start inhaling through device immediately prior to acutation | Actuation prior to inhalation | X |  |  |  |  | Potential for reduced dose |  | 4
| c. |  | Actuation after a second or more of inhalation | X |  |  |  |  | Potential for reduced dose, in particular to small / alveolated airways |  | 4
| d. | Actuate canister fully | Canister not actuated | X |  |  |  |  | No dose |  | 10
| e. |  | Canister only partially actuated | X |  |  |  |  | Potential for severely reduced dose |  | 7
| f. | Continue to inhale following actuation | Inhalation stopped on or soon after actuation | X |  |  |  |  | No dose / potential for severely reduced dose |  | 10 / 7
| g. | Achives target inhalation flowrate | Inhalation flowrate at too fast. | X |  |  |  |  | Potential for severely reduced dose | Can be a large effect for patients with small thoats. | 7
| h. | Release the canister with the pMDI still orientated in the valve down position | Incorrect orientation for valve re-filling stage (canister release). | X |  |  |  |  | Potential for severely reduced dose |  | 7
| i. | Actuate no more than a single unit dose for each inhalation | Multiple unit doses actuated during one inhalation | X |  |  |  |  | Potential for severely increased dose for that inhalation |  | 7
3.2.3 Inhale Loaded Unit Dose (BAMDI)
| a. | Inhale at high enough flowrate actuate BAM | Inhalation flowrate too slow to actuate BAM |  | X |  |  |  | No dose |  | 10
| b. | Inhale at sufficiently low flowrate to avoid throat deposition and deliver a good lung dose. | Inhalation flowrate at too fast. |  | X |  |  |  | Potential for severely reduced dose | Can be a large effect for patients with small thoats. | 7
| c. | Continue to inhale following actuation | Inhalation stopped on or soon after actuation |  | X |  |  |  | No dose / potential for severely reduced dose |  | 10 / 7
| d. | Actuate mechanism to refill subsequent unit dose (close cap or lower lever) within specified time | Mechanism for subsequent unit dose refill not actuated within specified time |  | X |  |  |  | Potential for severely reduced dose | Important for suspensions | 7
| e. | Ensure valve is in correct orientation for subsequent unit dose refill. | Incorrect valve orientation while actuating mechanism for filling of subequent unit dose (lever/cap closing). |  | X |  |  |  | Potential for severely reduced dose |  | 7
| f | Close cap/return dose lever to re-set dosing mechanism between actuations (if required) | Dosing mechanism not re-set if do not close cap so user receives no dose on subsequent inhalations |  | X |  |  |  | Potential for no dose on subsequent dose |  | 10
3.2.4 Inhale Loaded Unit Dose (SMI)
| a. | Start inhaling through device immediately prior to acutation | Actuation prior to inhalation |  |  | X |  |  | Potential for reduced dose |  | 4
| b. |  | Actuation after a second or more of inhalation |  |  | X |  |  | Potential for reduced dose, in particular to small / alveolated airways |  | 4
| c. | Press dose release button | Does not press dose release button |  |  | X |  |  | No dose |  | 10
| d. | Continue to inhale following actuation | Inhalation stopped on or soon after actuation |  |  | X |  |  | No dose / potential for severely reduced dose |  | 10 / 7
| e. | Achives target inhalation flowrate | Inhalation flowrate at too fast. |  |  | X |  |  | Potential for reduced dose | Reduced impact for SMI due to small droplet size, but can affect subjects with small throats. | 4
3.2.5 Inhale Loaded Unit Dose (Manually Loaded DPIs)
| a. | Achives target inhalation flowrate to deliver powder dose. | Inhalation flowrate too slow |  |  |  | X |  | Potential for severely reduced dose | Device dependent. | 7
| b. |  | Inhalation flowrate too slow at start of inhalation |  |  |  | X |  | Potential for reduced dose | May see reduced fine particle mass. Product dependent. | 4
| c. | Inhale sufficient air volume to deliver complete dose. | Incomplete or shallow inhalation |  |  |  | X |  | Potential for severely reduced dose |  | 7
| d. | Takes second inhalation if specified for product (to ensure delivery of residual dose remaining in capsule). | Second inhalation not taken (where specified for product). |  |  |  | X |  | Potential for reduced dose | Tobramycin requires second inhalation of each capsule. | 4
3.2.6 Inhale Loaded Unit Dose (Pre-Loaded DPIs)
| a. | Achives target inhalation flowrate to deliver powder dose. | Inhalation flowrate too slow |  |  |  |  | X | Potential for severely reduced dose | Device dependent. | 7
| b. |  | Inhalation flowrate too slow at start of inhalation |  |  |  |  | X | Potential for reduced dose | May see reduced fine particle mass. Product dependent. | 4
| c. | Inhale sufficient air volume to deliver complete dose. | Incomplete or shallow inhalation |  |  |  |  | X | Potential for severely reduced dose |  | 7
| d. | Takes second inhalation if specified for product (to ensure delivery of residual dose remaining in capsule). | Second inhalation not taken (where specified for product). |  |  |  |  | X | Potential for reduced dose | Tobramycin requires second inhalation of each capsule. | 4
| e. |  | Second inhalation taken after acuating the device a second time (resulting in two doses where one was intended) |  |  |  |  | X | Strong likelihood of overdose | Patient would almost certainly receive 2x intended dose. | 10
3.3 After Inhaling Dose
3.3.1 After Inhaling Dose (All Inhalers)
| a. | Take correct number of unit dose capsules/ blisters/ actuations as needed for correct therapeutic dose | Correct number of unit dose capsules/blisters/actuations not taken. | X | X | X | X | X | Potential for severely reduced dose |  | 7
| b. |  | Subsequent doses taken without preparing a new unit dose | X | X | X | X | X | Potential for severely reduced dose |  | 7
| c. | Rinse mouth and swallow or expectorate (where specified) | Mouth rinse not completed | X | X | X | X | X | No impact | Increased potential for side effects in oropharynx depending on drug safety profile (e.g. Inhaled Cortico-Steroids) | Not Scored, should be scored by project team.
| d. | Replace cap after use | Cap not replaced after use | X | X | X | X | X | No impact | Infection hazard | Not Scored, should be scored by project team.
3.3.2 After Inhaling Dose (Manually-Loaded DPIs)
| a. | Unload used capsule/blister and check for good emptying (manually-loaded devices). | Capsule/blister not removed or checked to ensure good emptying |  |  |  | X |  | Potential for reduced dose | Unit dose DPIs only | 4